## Slide 1
A Conformal Intervertebral Disc Prosthesis
On the acquiring and modeling of the geometry
Edwin H.E. Bennink, Bart M. ter Haar Romeny
Eindhoven University of Technology
Nico F. Kamperman
Netherlands Organization for Applied Research (TNO)
eindhoven
university
of technology

## Slide 2: Why do we need a conformal disc prosthesis?
About 70-80% of all people experience a lower back complaint in a particular period of their lifetime;
Within 20% of this group, some kind of arthrosis arises which may lead to chronic pain;
It is generally accepted that degeneration of the intervertebral disc plays an important role in the causes of these complaints.

## Slide 3: Why do we need a conformal disc prosthesis?
Fixation of joints leads to overload of adjacent joints;
Great successes with hip and knee joint replacements.

## Slide 4: Why do we need a conformal disc prosthesis?
Disc prosthesis can not be fixed into the vertebrae; it has to be fixed onto its surface;
Geometric conformance is a welcome addition to today’s disc prostheses.
Today’s disc prosthesis,
Alpha Klinik, Munich

## Slide 5: Anatomy of the vertebrae and the disc
vertebral body
intervertebral disc
facet joint
neural arch
sagital view
transversal view
50 mm
35 mm

## Slide 6: Anatomy of the vertebrae and the disc
flexible joint
conformal endplate
vertebral body
prosthesis

## Slide 7: What to do?
Patient
Scan
CAD Model
Prosthesis
CT
Segmentation
High speed
milling
Segmentation of the
intervertebral volume
DICOM data
From CT scan
STL file

## Slide 8: Why use Mathematica?
Mathematica is capable of importing DICOM and exporting STL;
MathVisionTools package contains various image analysis functions;
The Mathematica programming language delivers unrivalled speed in prototyping.

## Slide 9: What kind of algorithm to develop?
Active shape algorithms have the features we want:
Continuous closed surface;
Robust;
Relatively easy to understand;
Triangulated surface can be exported to STL.

## Slide 10: How does such an algorithm work?
The ‘balloon’ approach of the algorithm uses pressure and stiffness forces and is guided by the image gradient field.
pressure vector
elasticity vectors
image gradient vector
image gradient field

## Slide 11: How does such an algorithm work?
Pressure: Fp(n) = cp Nn
Elasticity: Fe(n) = ce (rn-1 + rn+1 − 2rn)
  Surface only: Fe(n) – (Nn·Fe(n)) Fe(n)
Image gradient: Fg(n) = cg G(rn)
Where Nn is the outer normal of point n,
rn, rn-1 and rn+1 are the position vectors of point n and its neighbours,
G(rn) is the gradient vector field,
and cp, ce and cg are constants.

## Slide 12: Segmentation in 3D

## Slide 13: Segmentation in 3D

## Slide 14: Improvements on the basic algorithm
Changes to the basic active shape algorithm:
Shape refines when growing larger;
Points which are not close to the edge are rejected from it;
Pressure force does not act on points close to the edge.
pressure vector
image gradient vector
new point
intervertebral
disc volume
cortical bone

## Slide 15: Segmentation in 3D
x-direction
y-direction
z-direction
information

## Slide 16: Segmentation in 3D
x-direction
y-direction
z-direction
information

## Slide 17: Adding more detail
Blur the image less and add new points to the shape to increase detail;
Smallest scale is σ = 0.63 pixels (5% error), which means the distance between the points is about 0.5 millimeter;
Noise is not a problem, because the shape is already on the edge.

## Slide 18: Adding more detail
x-direction
y-direction
z-direction
information

## Slide 19: Adding more detail
x-direction
y-direction
z-direction
information

## Slide 20: Future research
TNO is validating the accuracy of this algorithm by measuring the surfaces of vertebrae of pigs;
There are parallel research projects on other aspects of the prosthesis going on, like the flexible joint (a kind of fiber strengthened hydrogel) and the endplate material and biocompatible coating;
TNO is planning a project to introduce this prototype in clinical practice.
flexible joint
conformal endplate
vertebral body